Bryant ConsultantsConcerned About Preparing   
Your Equipment For A Temporary Practice Closure?

**Here is a helpful checklist to help guide you through the process.  Special thanks to Benco Dental for the information:**

**Your service team can help your offices to restart the equipment when it is time to do so.**

* **WATER BOTTLE:** Remove, empty, replace, run system to purge water from unit lines; remove water bottle, let air dry
* **AIR COMPRESSOR**: Power down and release all air; drain tanks where applicable
* **VAC LINES**: Flush all vac lines; replace sediment filters and all disposable traps on units
* **AMALGRAM SEPARATORS**: Replace amalgam separators due to prophy paste hardening
* **DRY VAC**: Power down and…
* **Midmark Power and G Vac**: If cutting main power, add WD-40 to pump drive piston to prevent lockup. If leaving power on and newer style system will autorun every so often to prevent lockup.
* **RamVacs:** Remove air filter to release pressure. Can hydro lock from oil drip? This is normal as motor overloads may trip on startup. Kill voltage to system.
* **Vac Systems**: Clean out tanks where applicable.
* **Mohave Vacs**: Leave breakers on or MMC can risk serious errors
* **WET VACS**: Turn off water supply after flushing lines and turn off the power
* **AUTOCLAVES**: Clean and drain; do not replace water; let air dry
* **HYDRIM**: Remove soap/salt and solution and run shipping cycle
* **DISTILLERS**: Empty all distillers and let dry
* **WATER PURIFICATION SYSTEMS**: Shut down water and purge with air until operatories and all faucets run dry
* **CHAIRS**: Leave chairs in upright position as high as they will go
* **HANDPIECES**: Oil all handpieces; clean and lube all HVE and SE handles
* **PLASTER TRAPS**: Remove and discard
* **COMPUTER SYSTEMS**: Complete backup on removable hard drive; take to a safe location offsite; power down all computers, servers, switches, etc.
* **ULTRASONICS**: Drain and clean; let air dry
* **WATER COLLECTION CANISTERS**: Empty all on cassette sterilizers
* **MODEL TRIMMERS**: Run water for a period of 10 minutes with unit left on to clean wheel as well as flush all drain lines
* **DUST COLLECTOR FILTERS**: Replace all in lab areas
* **AIRLINES**: Bleed all airlines in operatories and lab areas
* **MILLS**: Thoroughly clean; drain coolant water
* **3D PRINTERS**: Remove all resin in machine and clean

**Contact Bryant Consultants if you have questions or need assistance:** 

(877) 768-4799

4889 Eagle Pointe Dr

[info@bryantconsultants.com](mailto:info@bryantconsultants.com)

[bryantconsultants.com](https://www.bryantconsultants.com/)